## Tool list
Measurement Categories | Tool Category | Company | Region | Tool Name | Description |  |  | Consider/not consider to have | Cost | Maturity | Area Requirement | Technical specifications | [9 empty] | Other information | Source
[12 empty] | Measurement time | Time between measurements | Warm up time | Accuracy | Repeatability | Sample size and geometry | Sample temperature | Operating temperature | Wavelength band | Operational Procedure
Solar-Weigthed Specular Reflectance | Reflectometers | Surface Optics | San Diego, CA U.S. | 410 Solar | Battery operated portable reflectometer. Ideal tool for field solar reflectance, absorbtance measurements and mirror assessment. Collects solar reflectivity data to calculate solar reflectance index (SRI) values. |  |  | Consider to Upgrade | $34086 (410-Solar measurement head+handheld command module+ two calibration coupon) | Commercial | Handheld instrument | 10 secs | 2 secs | 90 secs | ±2% | ±.001 units | Flat: ≥0.5 in dia; Radius of curvature; 6 in convex; 12 in concave | Ambient or heated/cooled to 0-100 °C | 0-40 °C | 335-380 nm, 400-540 nm, 480-600 nm, 590-720 nm, 700-1100 nm, 1000-1700 nm, 1700-2500 nm | 1. Calibration with a specular calibration coupon and with a diffuse calibration coupon 2. Hold the device against the surface to be tested. 3. Press the trigger on the handle for measurements. 4. Results will display on the touch screen. Measurements can be taken on curved surfaces. Re-calibration is needed every 30 minutes. |  | https://surfaceoptics.com/products/reflectometers-emissometers/solar-absorptance-measurements-410/
|  |  |  | ET100 Emissometer | Measures directional reflectance at six bands in the thermal infrared spectrial region at two incidence angles (20° and 60°). |  |  | Consider to Upgrade | $41952 (ET100 measurement head+handheld command module+calibration coupon) | Commercial | Handheld instrument | 10 secs | 2 secs | 90 secs | ±3% | ±.005 units | Flat: ≥0.5 in dia; Radius of curvature; 6 in convex; 12 in concave | Ambient or heated/cooled to 0-100 °C | 0-40 °C | 1500-2000 nm, 2000-3500 nm, 3000-4000nm, 4000-5000 nm, 5000-10500 nm, 10500-21000 nm | 1. Calibration with a specular gold coupon measuredin reference to a coupon measured by NIST. 2. Hold the device against the surface to be tested. 3. Press the trigger on the handle for measurements. 4. Results will display on the touch screen. Measurements can be taken on curved surfaces. Re-calibration is needed every 30 minutes. |  | https://surfaceoptics.com/products/reflectometers-emissometers/et100-thermal-hand-held-emissometer/
|  |  |  | SOC-410 DHR | Directional Hemispherical Reflectometer is the first of its class handheld reflectometer for quick and accurate measurements of optical constants of surfaces in the infrared spectral region. It is the only fully portable hand-held device for quick inspection of coatings and materials in the field. The DHR head measures the integrated surface reflectance at two incidence angles (20° and 60°) and the optical constants of materials for six spectral regions. |  |  | Not consider to have | $57570 (410-DHR measurement head+handheld command module+calibration coupon) | Commercial | Handheld instrument | 10 secs | 2 secs | 90 secs | ±3% | ±.005 units | Any surface; convex 6 in radius; concave 12 in radius | Ambient or heated/cooled to 0-100 °C | 0-40 °C | 900-1100 nm, 1900-2400 nm, 3000-4000nm, 3000-5000 nm, 4000-5000 nm, 8000-12000 nm | 1. Calibration with a specular gold coupon 2. Hold the device against the surface to be tested. 3. Press the trigger on the handle for measurements. 4. Results will display on the touch screen. Measurements can be taken on curved surfaces. Re-calibration is needed every 30 minutes. |  | https://surfaceoptics.com/products/reflectometers-emissometers/soc410-dhr/
|  | Abengoa | Spain | Condor | Measures reflectance with a high level of reliability in six wavelengths and incidence angle of 12°, characterizing the response of a mirror at different wavelengths along the sun spectrum. Has several optical channels made up of LED-type light source and a photodiode, which detects the radiation reflected in the central wavelength of the LED. |  |  | Likely not consider | 25720 | Commercial | Handheld instrument | 10 secs |  | 5-10 minutes | ±0.2% (3σ) | ±0.001 | Both flat and curved surfaces up to 4mm thickness |  | -20 - 55 °C | 435 nm, 525 nm, 650 nm, 780 nm, 940 nm, 1050 nm | 1. Calibration (Autocalibrated with an accessory pattern mirror) 2. Hold the device against the surface to be tested, the measurement system will be activated. 3. The instrument can operate in automatic mode or manually. Measurements can be taken on curved surfaces. Re-calibration is needed when predetermined temperature range is exceeded. | 1. No influence of the curvature 2. Influence of temperature is detectedby the instrument. 3. No influence of the ambient light. 4. No dependence of the operator. | https://aragonphotonics.com/condor-solar-reflectometer/
|  | Device & Services Co. | Dallas, TX U.S. | 15R-USB | Designed for field and lab measurements of flat or curved specular reflectors. Powered by internal rechargeable Nicad batteries. Has USB port for maintaining and downloading datasets. Ideal for monitoring soiling of solar reflectors in field. It measures the specular reflectance at incidence angle 15° |  |  | Already have | 19000 | Commercial | 12.7cm dia *16cm height | Take readings 2-3 per minute |  |  | ±0.2% (at 660nm) | ±.002 units |  |  | 0-50 °C | 635-685nm, Peak at 660nm | 1. Position the device against the reflector surface. 2.Adjust the support for the maximum reading shown on the display. Measurements can be taken on curved surfaces. Re-calibration is needed periodically. | 1. No influence of the curvature. 2. No influence of the temperature. 3. No influenceof the ambient light. 4. High dependence of the operator. 5. The type-B uncertainty for most reflector types are 0.006 | https://www.devicesandservices.com/prod02.htm https://www.researchgate.net/publication/339462689_Uncertainty_Study_of_Reflectance_Measurements_for_Concentrating_Solar_Reflectors
|  |  |  | 15R-RGB | Adds the capability to measure reflectance in five broad wavelength bands from blue to near IR using a white LED source and selectable filters. |  |  | Already have | 23500 | Commercial | 12.7cm dia *15cm height | Take readings 2-6 per minute |  |  |  | ±.002 units |  |  | 0-50 °C | Band red, green, blue, white and IR filters: 460 (±50), 550 (±50), 650 (-40 +150), 720 (-40 +100) | 1.Select the filter and the aperture. 2. Position the device against the reflector surface. 3.Adjust the support for the maximum reading shown on the display. Measurements can be taken on curved surfaces. Re-calibration is needed periodically. |  | https://www.devicesandservices.com/prod04.htm
|  |  |  | SSR 6 | Features accurate solar reflectance measurements for a variety of global and beam normal irradiance models. The radiation reflected by the sample is measured at incidence angle 20° |  |  | Likely not consider | 41500 | Commercial | 51cm*34cm*20cm | One measurement every 8 secs |  | 5-10 minutes | ±0.2% | ±.003 units | Minimum sample size:1.1 inch in diameter. 1.65 inches across |  | Max 60 °C | Band UV, blue, red, and IR: 300-2500nm | 1. Calibration: The instrument is calibrated with one of the three standards and the calibration is accomplished with a signle key stroke. 2. Placing the sample over the measurement port, and leave the sample during the whole on and off cycle. A tungsten lamp will show the measurement progress by on and off. The new results will display when the lamp is turned off. |  | https://www.devicesandservices.com/prod01.htm
|  |  |  | AE1 RD1 | Special purpose instrument for measuring thermal emittance with the sample at ambient temperature. |  |  | Likely not consider | 4550 | Commercial | 15.2cm*10.5cm*15.3cm | 1 minutes and a half |  | 30 minutes |  | +/- 0.01 emittance units | Minimum flat surface areas: 1.5" in diameter | 130F maximum |  |  | 1. Check the offset on the Voltmeter by inserting a shorting plug and see if it display 0. 2. Calibration by two standard surfaces(one high emittance standard, one low emittance standard) 3. Put the sample on the heat sink. 4. Let the emissometer resting on the high emittance standard sample. Adjust the RD1 to indicate the emittance. 5. Place the emissometer on the sample and read emittance directly on the RD1. |  | https://www.devicesandservices.com/prod03.htm
|  | CSP Services GmbH | Germany | TraCS | Measures the effective reflectance including cleanliness of a sample solar mirror continuously. It compares between measured beam irradiance (DNI) and reflected beam irradiance from the sample mirror. |  |  | Reach out to Michael and Giovanni | 8200 | Commercial | Operating area: R0.85m *R0.625m * R0.7m | Average: 10 minutes |  |  | ±1.8% | ±.01 units | Limited adaptable to mirror curvature |  | -20-50 °C | Full solar spectrum | 1.Calibration: Cleanliness: sample mirror in clean state; Reflectance : reference mirror | • The TraCS is designed for Solys2 tracker only. It is not compatible with any other sun tracker model currently • With the mounting arm, the tracker needs an increased free area around it to operate. Please see attached drawing • With the TraCS installed, the tracker top mounting plate must be replaced with a customized top mounting plate that does not prevent rotation of the TraCS mounting arm. This new top mounting plate is included in the delivery, but it is important to know that the tracker then only offers 2 installation places for pyranometers/pyrgeometers (instead of 3) • You will need a 24V DC power supply outlet for motor power supply | https://www.cspservices.de/wp-content/uploads/CSPS-TraCS-Soiling.pdf
|  | PSE AG | Germany | PFlex | Portable Soiling Handheld reflectometer. Uses a mobile device for user interface. Offers precise measurements as well as data storage and transfer. |  |  | Not consider | It was not manufacturered any more | Commercial | Handheld instrument | 8 secs |  |  | ±1% | ±.005 units |  |  | -15-60 °C | 470nm(±25), 525nm(±25),625nm(±25) | 1. Calibration: Using a reference mirror, and a black reference for offset |  | https://www.pse.de/test-equipment/thermal-collectors/sensors/pflex-portable-handheld-reflectometer/
| Spectrophotometer | PerkinElmer | U.S. | LAMBDA 1050 | LAMBDA 1050 has a triple detector capability for ultra-high UV/Vis/NIR performance for wavelengths up to 3300 nm with extreme sensitivity in the NIR region (800-2500 nm) and higher energy throughput. For applications such as highly reflective and anti-reflective coatings, all types of glass from clear to highly absorbing safety glass, optical filters of all types from the deep UV through the NIR and many more applications requiring the best photo- dynamic range and the widest possible sampling capability |  |  | Already have | ~$122000 Can add another Laser alignment accessory (TAMS )~$100000 | Commercial | 1.02m*0.74m*0.3m | Depend on wavelength range and interval |  | ~20 minutes | Wavelength accuracy: UV/Vis: ±.080 nm; NIR: ±.300 nm Photometric accuracy: by different method : ±.0003~±.008 A | Wavelength: UV/Vis: ≤.010 nm; NIR: ≤.040 nm | Primary sample compartment: 200mm x 300mm x 220mm. Secondary sample compartment: 480mm x 300mm x 220mm |  | 10-35 °C | 175 - 3300 nm (N₂ purge required below 185 nm) | 1. Flip the toggle switch to turn on the device. 2. Turn on the computer, and open and login to the Perkin Elmer UV WinLab software. 3. Select the instrument method, input the sample information. 4. Use the Data Collection icon to change data collection parameters. 5. Remove any old samples left in the instrument before as needed. 6. Open the lid and insert the sample. 7. Start the test. 8. Save the results. |  | https://camcor.uoregon.edu/site/wp-content/uploads/2012/11/Lambda-1050-spec.pdf https://www2.chemistry.msu.edu/faculty/mccusker/SOPs/SOP-Lamda1050.pdf
|  | Konica Minolta | Japan/U.S. | CM-700d/600d | Sphectrophotometer is a handheld, portable measurement instrument designed to evaluate the color and appearcance of samples very small to large in size, including objects with flat, shape, or curved surfaces. Measuring incidence angle is 8° |  |  | Not Consider | ~$9940 | Commercial | Handheld instrument | 1 sec | 2 secs |  |  | ±.001 units |  |  | 5-40 °C | 400-700nm with 10nm steps | 1. Turn power on. 2.Registering conditions. 3. Attach/remove a target mask if needed; Zero Calibration if needed. 4.White calibration with the white calibration cap. 5. Take measurements by place the specimen measuring port on the sample. 6. Press the measuring button. 7. If checking the color difference, perform the color difference target color data operation before the measurement. No re-calibration is needed during 3-4 hours measurement. | For CM-700d: 1. No influenceof temperature. 2. No influence of ambient light. 3. Little dependence on operator. | https://sensing.konicaminolta.us/us/products/cm-700d-spectrophotometer/ Cost source: https://www.sinodevices.com/Minolta-CM-700d-Spectrophotometer.html
|  |  |  | CM-26d 26dG | Next generation of portable sphere spectrophotometer with horizontal alignment, a design convenient for measuring the color of flat or large samples. |  |  | Not consider | ~$13331-$14331 | Commercial | Handheld instrument | CM-26d: 1 sec; CM-26dG: 0.7 sec | CM-26d: 2 sec; CM-26dG: 1.5 sec |  |  | <0.02 |  |  | 5-40 °C | 360-740nm with 10nm steps | 1. Turn the power on; 2. Measurement condition setting; 3. Measurement area selection; 4. Zero Calibration if required; 5. White Calibration with a calibration stage; 6. Set the specimen meauring portal on the sample and press the measurement button; |  | https://sensing.konicaminolta.us/us/products/cm-26d-spectrophotometer/
| Dust Sampler | TSI | USA | DustTrak DRX Aerosol Monitor | It can simultaneously measure both mass and size fraction. DustTrak DRX monitors are multi-channel, battery-operated, data-logging, light-scattering laser photometers that give you real-time aerosol mass readings and collects a gravimetric sample. They use a sheath air system that isolates the aerosol in the optics chamber to keep the optics clean for improved reliability and low maintenance. |  |  | Consider to have | $12635 (With External Pump: $13545) | Commercial | Desktop size (0.135m*0.216m*0.224m) External Pump: 0.1m*0.18m*0.09m | Resolution: +/-0.1% of reading |  |  | Flow accuracy: Precision automatic flow control to within +/- 5% |  | Particle Size Range: 0.1-15μm | [5 empty] | https://eco-rentalsolutions.com/product/air-monitoring-equipment/dust-monitoring-equipment/dust-monitor-dust-trak-drx-tsi-8533/
|  | Thomson Environmental Systems | Australia | Dust Master Pro 7000 | Real-time measurement of up to 5 PM fractions simultaneously Capable of measuring PM10 and PM2.5 (the most common fractions for measuring dust that is harmful to health) as well as PM4, PM1 and Total PM |  |  | Consider to have, the dust sampler set up at NSTTF | 15878 | Commercial | System Dimension: 0.3m*0.2m*0.35m Heated inlet dimension: 0.5m*0.048m | Measurement resolution: 0.001mg/m^3 Flow rate: 5LPM |  |  | Flow accuracy: Precision automatic flow control to within +/- 1% |  | Particle Size Range: 0.2-18μm
Measurement Catgegories | Tool Category | Company | Region | Name | Description |  |  | Consider/not consider to have | Area Requirement | Maturity | Cost | Technical specifications | [7 empty] | Source
[12 empty] | Components |  | Measurement time |  | Measurement repeatability | Measurement accuracy | Measurement uncertainty
Opto-Mechanical Errors | Structured Light Reflection | DLR/CSP services | Germany | Qdec | Optical measurement system for control of the shape accuracy of solar reflector panels and concentrators. |  |  | Likely not, similar to SOFAST, but may be more expensive |  | Commercial |  | A camera: capture the reflection of a screen shown in the tested mirror facet. Projector: display images on the screen Computer: run the software in MATLAB |  | QDec Offlne < 30 s, evaluation time <40s; QDec Inline < 5 s, evaluation time<10s Qdec: 1min/heliostat Qdec-M-Helio: <5 mins (Measurement+evaluation) |  | Qdec-M-Helio: <0.1 mrad locally | Qdec-M-Helio: 0.2 mrad globally | Qdec-M-Helio: Local spot (RMS): <0.5 mrad; Global value (RMS): < 0.2 mrad |  | https://www.cspservices.de/wp-content/uploads/CSPS-QDec.pdf https://doi.org/10.1016/j.solener.2010.01.010 https://elib.dlr.de/125526/1/CSPS_Presentation_SolarPACES_181003_Research2Industry_still_logos.pdf
|  | Sandia | U.S. | SOFAST | Sandia Optical Fringe Analysis Slope Tool. Measures surface slope of CSP optics. It uses deflectometry to obtain a high-resolution map of mirror surface slope, and then uses this to generate a variety of analytic products characterizing optical error. |  |  | Plan to have, but need to look at the footprint | ~100m^2 | Available |  | A camera: capture the reflection of a screen shown in the tested mirror facet. Projector: display images on the screen Computer: run SOFAST software in MATLAB |  | Cycle time: ~10s |  |  |  | Twist measurements accurate to: ±.03 mrad/m; Local slope measurement accurate to: ±.05 mrad; Global slope measurement accurate to: ±.005 mrad [2] |  | [1] https://www.osti.gov/servlets/purl/1855054 [2] https://doi.org/10.1115/1.4024251 [3] https://www.osti.gov/servlets/purl/1106023
|  |  |  | AIMFAST | Alignment Implementation for Manufacturing using Fring Analysis Slope Technique. An extension of SOFAST that automates data collection, implements alignment strategies, and provides real time mirror angle corrections to actuators beneath the dish |  |  | Likely not recommend |  | Available |  | A camera: capture the reflection of a LCD screen shown in the tested mirror facet. Colored panels: help automatically locate the facets. Projector: display images on the screen Computer: run AIMFAST |  | Cycle time: 4-5s; With a full set of tools, the dish could be aligned in 11 minutes including tool installation and removal. |  |  | <0.25 mrad RMS on dish systems. | 2.2 mradRMS without centering; 0.09mrad after centering |  | [1] https://asmedigitalcollection.asme.org/solarenergyengineering/article/133/3/031018/456214/AIMFAST-An-Alignment-Tool-Based-on-Fringe [2]https://www.osti.gov/servlets/purl/1110492 [3] https://www.osti.gov/servlets/purl/1143874
| Canting |  |  | H-FACET | Uses a high resolution digital camera mounted on top of a receiver tower to observe the image of a startionary target reflected by a heliostat. Custom image processing sofrwate compares specific measurement points on the target reflection image with the cooresponding measurement points on an ideally reflecting heliostat. Deviations between these points reveal facet misalignmnets |  |  | Outdated, Not recommend |  | Not commercially available |  | A camera located at the tower top,an in-field target, computer for running H-FACET in MATLAB |  | Alignment time < 1 hour/heliostat [1] |  |  | Facet alignment accuracy: 0.3mrad RMS [2] Total system accuracy: 1.19 mrad on five heliostats. |  |  | [1]https://www.osti.gov/servlets/purl/1661779 [2]https://www.osti.gov/servlets/purl/1090214 [3] https://doi.org/10.1115/ES2012-91039
|  |  |  | UFACET | Universal Field Assessment, Correction, Enhancement Tool. Similar to H-FACET except camera is not fixed on the tower but mobile moving around the field. Using an unmanned aerial system (UAS) |  |  | Outdated, Not recommend |  | Under development |  | A mobile camera with an unmanned aerial system; In field target; Computer for running MATLAB code) |  | Within flight time, about 15-20 mins for many heliostats |  |  | Canting error estimates between UFACET and LIDAR: 0.8mrad-1.0 mrad RMS for both azimuth and elevation directions.[2] |  |  | [1]https://www.energy.gov/sites/default/files/2019/04/f62/CSP%20Summit2019%20SNL-NREL%20Yellowhair-Zhu%20UFACET-NIO.pdf [2] https://doi.org/10.1063/5.0087057
|  | NREL | U.S. | NIO | Non-Intrusive Optics Approach. Given a series of images of a heliostat compare the reflected tower edge in each facet with that of a reference facet with known canting. Distortions and misalignments can be measured to calculate optical errors |  |  | Already have | Operational area: ~10000m^2 Drone:~1m^2 | Under development |  | A high-resolution camera mounted on an automated aerial drone; | [5 empty] | Heliostat tracking error: 1.71 mrad (0.12 mrad at a confidence level of 95%); Hleiostat RMS of slope error: 1.22 mrad (0.05 mrad at a confidence level of 95%); Individual Mirror facet slope error:1.86 mrad-1.38 mrad( at 95% confidence level 0.22 mrad RMS) at given test condition in the study. [2] |  | [1]https://doi.org/10.1016/j.solener.2020.06.093 [2]https://www.osti.gov/servlets/purl/1677476
Measurement Catgegories | Tool Category | Company | Region | Name | Description |  |  | Consider/not consider to have | Cost | Maturity | Area Requirement | Sensitivity |  | Technical specification | [10 empty] | Source
Heliostat Shape | Laser Scanning | Sandia | U.S. | LiDAR (with FARO S70 Scanner) | Acquires highly accurate point cloud measurements across several heliostats revealing any optical errors on the mirror. |  |  | Need to consider the the price | Faro s70 cost: 150m range scanner: $47000-48000 with 20% discount; 350m range: $56000-57000 with 25% discount | LiDAR: Available | Scanner: 0.23m*0.183m*0.103m Canbe set as far as 70m, as close as 0.6m from the target | Canting angles are slightly influenced by the elevation angle. |  | Components | Measurement speed | Time | Measurement error | Range resolution | Repeatability | Accuracy | Range | Field of view | Operating temperature | Basic Operational Procedure | [1]https://www.osti.gov/servlets/purl/1784100 [2]https://www.faro.com/en/Products/Hardware/Focus-Laser-Scanners
[14 empty] | 3D Laser scanner; Computer(run software) | Up to 1 (mil.pts/sec) | Entire process(Scan+process) less than 5 mins per heliostat | ±1mm (at around 10m and 25m) | ±2mm | Average RMS error in Azimuth: 0.21mrad; in Elevation: 0.20 mrad. | 0.25 mrad error accuracy requirement is achieved; Angular accuracy: 19 arcsec for vertical/horizontal angles 3D Point Accuracy: 2mm at 10m, 3.5mm at 25m, add 0.1mm/m after 25m | 90% Reflectivity (White): 0.6-70m 10% Reflectivity (dark-gray): 0.6-70m 2% Reflectivity (black): 0.6-50m | 300° vertical, 360° horizontal | 5-40 °C | A software automatically controls the scanner and download the scan data, isolates theheliostat data from the full scan, filters the points associated with each individual facet and calculate the canting angles. Such software can measure heliostat canting errors to less than 0.25 mrad accuracy within 5 mins processing time per heliostat. [1]
|  | NREL | U.S. | ReTNA | NREL's Reflected Target Non-intrusive Assessment (ReTNA) system measures mirror surface slope using a series of images of a printed target pattern reflected on the mirror's surface. Photogrammetry and computer vision methods are used to automatically identify the target, mirror, and camera in a common reference frame, reducing the required precision of the setup. A moving camera and modular printed target reduce the size and cost of the installation. |  |  | Plan to have | No additional cost | Under Development | Equipment area: ~5m^2 | Uncertainty <0.2mrads | [12 empty] | Under Development
| Hartmann type methods | Sandia & NREL | U.S. | SHOT | A method for characterizing the surface figure and optical performance of point focus solar concentrators |  |  | Outdated, not recommended | [16 empty] | https://www.nrel.gov/docs/legosti/old/7296.pdf
|  |  |  | VSHOT | A tool for characterizing Large, imprecise reflectors |  |  | Outdated, not recommended |  |  |  | Combined offset/tilt: mirror moved 10.6 degree to the right and tilted 5.3 degree to centerthe return pattern.The RMS error was 260% larger for a total of 0.23 mrad. |  | Components |  |  | Measurement error | [7 empty] | https://www.osti.gov/servlets/purl/396613
[14 empty] | Computer system; Laser&scanner;Camera; Video board |  |  | ±.05mrad for typical solar reflectors;less than ± .2mrad for small diameter. | [6 empty] | 1. Video andlaser scanner calibration.
|  |  | Spain | SCCAN | The SCCAN technique relies in simple principles of geometrical optics. The image of an ideal paraboloidal mirror for a point source (object at infinity) on the optical axis is a geometrical point at the focus. Therefore an observer located at the focus would see the whole mirror surface “bright” (i.e. light comes from all mirror points). Because a realistic mirror is far from ideal, the image of a point source has a finite size. In fact the image of a point source out of the optical axis has a finite size even for an ideal mirror. |  |  | Likely not recommended | [16 empty] | https://www.sciencedirect.com/science/article/pii/S0038092X03002597?via%3Dihub
| Photogrammetry |  |  | Close Range Photogrammetry | Involves a network of multiple photographs of a targeted object or heliostat taken from a range of viewing positions to obtain 3D coordinate data for the target. There are a variety of methods for specific solar dishes or heliostats |  |  | Plan to have | no additional cost | [15 empty] | https://asmedigitalcollection.asme.org/solarenergyengineering/article/127/1/94/459229/Photogrammetry-A-Powerful-Tool-for-Geometric
Measurement Catgegories | Name | Company | Consider/not consider | Cost | Description | Measurement speed |  |  | Area Requirement | Uncertainty analysis |  |  | Source
Sunshape Measurement | Sun and Aureole Measurement System (SAM) | DLR, Visidyne |  | SAM+ Sun photometer ~90000$ | The measurement system consists of the SAM instrument, an AERONET sun photometer, and postprocessing software, which can estimate the broadband sunshape and get information on aerosols and solar spectrum based on radiative transfer models. | 5s per measurement |  |  |  | For different CSRs (circumsolar ratio), the uncertainty is 0.053*broadband CSR+0.013 |  |  | https://doi.org/10.1016/j.solener.2016.01.051 https://digitalcommons.usu.edu/cgi/viewcontent.cgi?referer=&httpsredir=1&article=1051&context=spacegrant
| Rotating Shadowband Irradiometer (RSI) Based Measurement System | CSP Services GmbH |  | ~14000$ | Conventional rotating shadowband irradiometer was used as a DNI measurement device, and without any hardware modification, it can enable the sunshape measurement from CSC (Circumsolar contribution) or CSR (circumsolar ratio). |  |  |  | 0.5m*0.1m*0.2m | Average standard uncertainties for corrected 10 min data of 2 % for global horizontal irradiance (GHI), and 2.9 % for DNI (for GHI and DNI over 300 W/m²) were found for the 2012 yearly dataset when separate GHI and DHI calibration constants were used. |  |  | https://uol.de/f/5/inst/physik/ag/enmet/download/fachtagung_2016/poster/Poster_CSPS_Enmet_2016_Bremerhaven.pdf https://doi.org/10.1063/1.5067201
| BPI-CSR 460 sensors | DLR&Black Photon Instruments |  |  | BPI-CSR460 uses two pyrheliometers with different opening angles and semiconductor based detectors, and the circumsolar irradiance, CSR and the sunshape can be measured from a specfic angular interval. | 0.1s per measurement |  |  |  | This results in minimal uncertainties of 0.015 for CSR < 0.05, 0.025 for 0.05 < CSR < 0.1 and 0.04 for higher CSR. |  |  | https://www.dlr.de/sf/en/desktopdefault.aspx/tabid-10436/12676_read-39752/ https://doi.org/10.1063/1.4822222
Measurement Catgegories | Tool Categories | Name | Company | Region | Description |  | Consider/not consider |  | Area Requirement | Cost | Technical Specification | [9 empty] | Source
[11 empty] | Measurement time | Repeatability | Absolute attenuation measurement accuracy | Power measurement accuracy | Measurement spectral range | Temporal and incident power stability deviation | Reference measurement distance | Distant measurement distance | Solar distribution standard
Atmosphere Attenuation Measurement | Based on optical spectrum analysis | ATMOS | ZEPREN Solutions | Spain | ATMOS is an atmospheric attenuation measurement system developed for CSP tower plants. It works autonomously and measures the atmospheric transmittance value every five minutes. The measurement is based on optical spectrum analysis. |  |  |  | Screen size: 2m*2m Reference device size: 1m*1m*1.5m Measurement device size: 2m*1m*1.5m | Price start from $100,000 and depend on the installation site. The system requires a customized installation for each type of plant | 60-120s; Integration time: 0.001-5s | <0.50% | 1% | <0.5% | 400-1650nm | <1% | 10-120m | 800-1500m | ISO 9050 | https://zepren.com/index.php/home/csp-instrumentation/ https://doi.org/10.1016/j.solener.2022.03.057
Atmosphere Extinction |  | Vaisala FS11 scatterometer+ CIMEL sun photometer/AERONET | Products: Vaisala, CIMEL System: DLR | Finland, France | The Vaisala FS11 scatterometer is based on a monochromatic near-infrared light source emission and measures the strength of scattering processes in a small air volume mainly caused by aerosol particles. To optimize the absorption and broadband correction procedure, additional sun photometer is used to enhance on-site atmospheric assumptions. |  |  |  | Vaisala FS11 : 2.8m*0.9m*1m Sun Photometer: Flycase box: 0.66*0.52*0.47m; Tripod box: 1.03*0.57*0.6m | Vaisala FS11: ~15000$ +CIMEL sun photometer: ~49000$ |  |  | For T1 km equal to 0.9 and a 10 min time resolution, an uncertainty analysis showed that an absolute uncertainty of about 0.038. For yearly averages several error influences average out and absolute uncertainties of 0.020
|  | long-path transmissometer Optec LPV4+ CIMEL sun photometer/AERONET | Products: OPTEC, CIMEL System:DLR | US, France | The Optec LPV4 long-path visibility transmissometer determines the monochromatic attenuation between a light-emitting diode (LED) light source at 532 nm and a receiver and therefore also accounts for absorption processes.To optimize the absorption and broadband correction procedure, additional sun photometer is used to enhance on-site atmospheric assumptions. |  |  |  | Optec LPV4 :0.45m*0.12m*0.15m Sun Photometer: Flycase box: 0.66*0.52*0.47m; Tripod box: 1.03*0.57*0.6m |  |  |  | For T1 km equal to 0.9 and a 10 min time resolution, an uncertainty analysis showed that an absolute uncertainty of about 0.057. For yearly averages several error influences average out and absolute uncertainties of 0.054.
Measurement Category | Tool Categories | Name | Company | Region | Description |  | Consider/not consider |  | Cost | Area Requirement | Response time | Sensitivity
Atmosphere attenuation &sunshape meaurement | Based on DNI | SOLYS2 Sun Tracker+CHP 1 Pyrheliometer | Kipp&Zonen | Netherlands | CHP 1 Pyrheliometer is attached to SOLYS2 Sun Tracker to monitor the direct normal irradiance. With a 5 degree field of view, the CHP1 pyrheliometer monitors the direct-beam sun irradiation. The CHP1 must always be facing the sun due to its narrow field of vision, so SOLYS2 Sun tracker rotates on two axes and uses a GPS receiver to keep the CHP1 always aimed at the sun. |  |  |  | USD 19845.9 | 0.5m*0.34m*0.38m | <5 s | Temperature Dependence of Sensitivity :<0.5%(-20 degree to +50 degree); Non-linearity:<0.2%; SOLYS2 Sun tracker pointing accuracy: < 0.1° passive tracking / < 0.02° active tracking (with optional sun sensor)
|  | Pyranometer CMP 10/21/22 | Kipp&Zonen | Netherlands | Compare to Pyrheliometer, Solar radiation measuring instruments with a hemispherical (180°) field of view are called pyranometers. It measures the global solar radiation |  |  |  | CMP 10: 2634$ CMP 21: 4141$ CMP 22: 8675$ |  |  | 7-14 μV/W/m² <+/-0.5% (-20 to 50 °C)
Measurement Category | Name | Company | Region | Description |  | Consider/not consider |  | Cost |  |  | Area Requirement | Response time | Sensitivity |  | Source
Wind Load Measurement | Strain gages(Model: Vishay LWK-06-W250D-350,-W250B-350) | Micro-Measurements | USA | Measuring strains |  | May not need to buy |  | W250B: $224.34 W250D: $510.6 |  |  | No area requirement |  | Resistance tolerance: +/-0.4% Strain range: +/-5% |  | https://micro-measurements.com/pca/detail/w250b
| Accelerometers (2460-10) | Silicon Designs | USA | A device that measures the vibration, or acceleration of motion, of a structure |  | May not need to buy |  | 1875 |  |  | No area requirement |  | Sensitivity: 400mV/g; Bias Calibration error: 0.25% of span (Typical);0.6% of span (Max) Output Noise: 23μg/(root HZ) |  | https://www.silicondesigns.com/_files/ugd/3fcdcf_c719ab27995c4606abeb93ece9eb916f.pdf
| Laser displacement sensor (OM30-L0350) | Baumer | USA | Measure mirror deflections |  | May not need to buy |  | 1500 |  |  | No area requirement |  | Repeat accuracy 2 … 30 µm Linearity error ± 0.1 % Mr , 50 ... 200 mm ± 0.15 % Mr , 50 ... 350 mm |  | https://www.baumer.com/us/en/om30-l0350-hv-yin/p/42847
| Dynamic Inclinometer (2GiG BH1-1800-0-2M) | 2GiG | USA | An instrument for measuring angles of slope (tilt), elevation or depression of an object with respect to the gravity while it is not stationary |  | Plan to have |  | 329 |  |  | No area requirement |  | Accuracy: 0.05° accuracy
| Sonic anemometers (WindMaster) | Gill | UK | The Gill WindMaster is a precision anemometer offering three-axis wind measurement data. This instrument will monitor wind speeds of 0-50m/s and provides sonic temperature, speed of sound and U, V & W vector outputs at 20Hz (32Hz optional). This anemometer is of aluminium/carbon fibre construction and is ideal for the understanding of turbulent flows, surface energy balance and scalar fluxes. |  |  |  | 3380 |  |  | R0.24m*0.75m height |  | Wind Speed accuracy: <1.5% RMS @12 m/s <1% RMS @12 m/s (Custom) Direction accuracy: 2° @12 m/s 0.5° @12 m/s (Custom) Sonic Temperature accuracy: -20°C to +30°C within ±2°C of ambient temperature Speed of Sound: < ±0.5% @ 20°C
| TSI-880 Automatic Total Sky Imager | Yankee Environment Systems | USA | An automatic, full-color sky imager system that provides real-time display of daytime sky conditions. In addition to present weather and sky monitoring, it provides and automated real-time panoramic image making it well suited for physical security applications. |  |  |  | ~$30,000 |  |  | 0.53m*0.48m*0.87m